setup()
Initialize system states to all offline, getStatus() will update them
 - chiller offline
 - humidity sensor offline
 - TECs offline
 - LCD offline
 - state UNKNOWN
Configure start/stop button and LED, fault/no-fault LED, and the rotary encoder
initSystem()
setInitialHumidityThreshold
 Take a humidity sample, set the threshold 10% above this reading
startChiller()
Always start the chiller on boot up
Derive the chiller status using get commands, set the chiller's running and on-line status
handleChillerStatus()
Derive the TECs status using get commands, set all the TEC's running and on-line status - all TECs are treated as a single unit, if one is not running, they are all not running, etc.
handleTECStatus()
Take a humidity reading, keep a running average over the past 6 seconds.
getHumidityLevel()
loop()
Every 10 seconds, update component status/health
 - Using get commands, get the chiller's
- status, running, set-point and circulating temperatures, on-line/off-line
- update the chiller's health statistics
 - Using get commands, get each TEC's
- status, running, set-point and object temperatures, on-line/offline
- update each TEC's health statistics
- Take a humidity sample
- keep a running average of samples over the past 6 seconds
Using the fetched/updated component information
 - update the on/off LED
 - fault/no-fault LED
 - update LCD display contents (see manageLCD() below for details)
 - shutdown the system due to any failure, i.e. chiller not on or chiller comm failure results in TECs being shutdown
 - humidity high event causes chiller and TECs to be shutdown
 - enable/disable on/off button and/or the humidity knob
getStatus()
handlMsgs()
Switches on box
Messages from the controlling PC
Set / Get
Chiller set point() or circulating temperature
Switch, knob, dial
operations
Get status
Set / Get
a TEC set point or object temperature
Turn on / off TECs
Get / Set Threshold
Get Humidity
1. Get humidity threshold
2. Set humidity threshold
3. Get current humidity
1. Set a TEC set point
 - use meerstetter commands
- return response : ack / nack
2. Get TEC set piont
 - return last temperature from getStatus() above
3. get a TEC's object temperature
1. Turn on all TECs
 - use meerstetter commands to turn on all TECs
 - if at most 1 TEC fails to turn on return NACK [ and shutDown() ?]
 2. Turn off all TECs
 - use meerstetter commands to turn off all TECs
 - if at most 1 TEC fails to turn off return NACK, else return ACK
1. Set chiller set point
 - use the Huber General Command to set the set point
 - return response : ack / nack
2. Get chiller set point
 - return last SetPoint from getStatus() above
3. Get the chiller circulating temperature
Switches/knobs/dials on box
1. on / off:
 - on, do startUp
 - off, do shutDown
2. adjust humidity threshold
In borh cases update the LCD as needed
Return sub-system status – these were gotten via getStatus() above
 - humidity alert present
 - all TECs running
 - chiller on-line
Processing continues with manageLCD()
 - There are 5 categories of messages displayed on the LCD, each category gets 3 seconds of display time.
 - If there is a high humidity or humidity sensor failure, the message “***HUMIDITY ALERT*** is constantly displayed. When the failure is cleared, normal operation is resumed.
 - Normal operation rotates through these messages. ‘system status’ column is always displayed. The ‘fail reason’ column is displayed only while there is ‘some thing’ bad. For example, if the chiller becomes unreachable, ‘system status’ would be ‘SYS shutdown’, w/ fail reason, ‘SYS FAILURE’ and the chiller would be ‘CHL STOPPED’ and CHL COM FAILURE’ .
Category system status fail reasons (only when there is a failure)
----------------------------------------------------------------------------------------------------------------------------------------------------
Overall System SYS initializing SYS START FAIL – some startup failure
        SYS ready SYS FAILURE - some run time failure
        SYS running
        SYS shutdown
TEC TEC stop TEC COM FAILURE
        TEC run
        (these include each TEC’s set point – only shows 3 TECs)
Chiller CHL running CHL COMM FAILURE
        CHL stopped
        (these include set point and current temperature)
Humidity Humidity X Threshold Y *** HUMIDITY ALERT ***
        SENSOR FAILURE
manageLCD()