setup()
Initialize system stats to all offline, system boot, startUpCmd, and getStatus() will update them
 - chiller offline
 - sensor offline
 - TECs offline
 - LCD offline
initSystemStates()
 - Startup the SHT sensor
 - take a humidity sample
 - if the sensor does not start or the humidity is too high shutDown() the system
 - shutDown updates the system running status and LCD face ( see manageLCD, if the humidity is too high or the SHT sensor is not functioning, the LCD manager keeps the message “Humidity Failure” on the LCD)
takeHumiditySample()
OK ?
shutDown()
(see below)
no
yes
 - There are 5 categories of messages displayed on the LCD, each category gets 3 seconds of display time.
 - If there is a high humidity or humidity sensor failure, the message “***HUMIDITY ALERT*** is contantly displayed. When the failure is cleared, normal operation is resumed.
 - Normal operation rotates through these messages. ‘system status’ column is always displayed. The ‘fail reason’ column is displayed only while there is ‘some thing’ bad. For example, if the chiller becomes unreachable, ‘system status’ would be ‘SYS shutdown’, w/ fail reason, ‘SYS FAILURE’ and the chiller would be ‘CHL STOPPED’ and CHL COM FAILURE’ .
Category system status fail reasons (only when there is a failure)
----------------------------------------------------------------------------------------------------------------------------------------------------
Overall System SYS initializing SYS START FAIL – some startup failure
        SYS ready SYS FAILURE - some run time failure
        SYS running
        SYS shutdown
TEC TEC stop TEC COM FAILURE
        TEC run
        (these include each TEC’s set point – only shows 3 TECs)
Chiller CHL running CHL COMM FAILURE
        CHL stopped
        (these include set point and current temperature)
Humidity Humidity X Threshold Y *** HUMIDITY ALERT ***
        SENSOR FAILURE
manageLCD()
Verify communication with all sub-systems
Initialize the SHT Sensor
 - sht.init()
Start the LCD
 - lcd.begin(), lcd.clear()
 - lcd.print “deftDevise Firmware ver 1.0”
Communicate with the chiller
 - do Verify command – get name
 - do Limit command – get limits
 - do General command – get status
Communicate with all TECs (fail if only one fails)
- do MeCom_COM_DeviceStatus
allDevicesPresent()
Stop the chiller
- do General Command w/o ‘O’ for temp control
Stop all TECs
 - MeCom_TEC_Oth_LiveEnable - disable LiveEnable
 - MeCom_TEC_Ope_OutputStageEnable – do static off
shutDown()
All pass ?
NO
startChiller()
Start OK ?
NO
loop()
manageLCD()
See above
 - paint all sub-system status, 3 seconds each, or show humidity alert persistently
Check the following
- store fetched/derived sub-system status, will return them in the get status message from control PC, see below
– Update the LCD status for each sub-system, used by manageLCD to show the status
Chiller running:
 - Send General Command to fetch all attributes
 - if ‘O’ != temperature control status, assume chiller is running, else fail and update the LCD w/ chiller fail message and return (leave this message on the LCD)
All TECs running:
 - Send MeCom_COM_DeviceStatus for all TECs
 - looking for 1 (Ready) or 2(Running), else fail and update the LCD w/ TEC failure message and return (leave this message on the LCD)
Humidity check
 - query the SHT Sensor
 - verify humidity is below the set threshold, else fail and update the LCD w/ humidity failure message and return (leave this message on the LCD)
getStatus()
Status OK ?
shutDown()
NO
Stop the chiller and the TECs
Stop chiller:
- do General Command w/o ‘O’ for temp control
Stop all TECs:
 - MeCom_TEC_Oth_LiveEnable - disable LiveEnable
 - MeCom_TEC_Ope_OutputStageEnable – do static off
 - do not update the LCD
Update sub-status, manageLCD() will use the sub-system status to reflect on the LCD
handlMsgs()
Switches on box
Messages from the controlling PC
Switch, knob, dial
operations
Get status
Set / Get
Chiller set point()
Set / Get
a TEC set point
Turn on / off TECs
Get / Set Threshold
Get Humidity
1. Get humidity threshold
2. Set humidity threshold
3. Get current humidity
1. Set a TEC set point
 - use meerstetter commands
- return response : ack / nack
2. Get TEC set piont
 - return last temperature from getStatus() above
1. Turn on all TECs
 - use meerstetter commands to turn on all TECs
 - if at most 1 TEC fails to turn on return NACK [ and shutDown() ?]
 2. Turn off all TECs
 - use meerstetter commands to turn off all TECs
 - if at most 1 TEC fails to turn off return NACK, else return ACK
1. Set chiller set point
 - use the Huber General Command to set the set point
 - return response : ack / nack
2. Get chiller set point
 - return last SetPoint from getStatus() above
Switches/knobs/dials on box
1. on / off:
 - on, do startUp
 - off, do shutDown
2. adjust humidity threshold
In borh cases update the LCD as needed
Return sub-system status – these were gotten via getStatus() above
 - humidity alert present
 - all TECs running
 - chiller on-line
 Code goes back to top of loop()